## Slide 1
JÖNKÖPING UNIVERSITY
PHASE 8.6.4: VALIDATION TEST V2
Test Cases — 8.6.4 — Tension test
Results & Analysis
Specimen: 3D-printed PLA, gauge length 80 mm. Markers on thin gauge section. Six measurement cycles at +1.0 mm crosshead travel (tension), preceded by C0 conditioning. Performed before button fix; sign verified via V3.
Run | Travel | Wait before | Preload at start | Note
T0 | +1.0 mm | (after C0) | +188 N residual | Includes seating transient
T3 | +1.0 mm | ~6.5 min | Re-tared | Early shakedown
T4 | +1.0 mm | ~2.5 min | Retained | Mid shakedown
T5 | +1.0 mm | ~2.0 min | Retained | Near steady-state
T6 | +1.0 mm | ~11 min | Retained | Steady-state (longer rest)
T7 | +1.0 mm | ~4 min | Retained | Final steady-state
Metric | T5 (steady) | T6 (steady) | T7 (steady) | ¦T6 vs T7¦ deviation
Peak force | +756.4 N | +758.1 N | +755.9 N | 0.29 % ✓
Peak stress | +9.45 MPa | +9.48 MPa | +9.45 MPa | 0.29 % ✓
Motor_Strain at peak | +0.01250 | +0.01250 | +0.01249 | <0.01 % ✓
ΔL_px | +1.6 px | +1.6 px | +1.6 px | 0.0 % ✓
Peak ε_c | +0.000958 | +0.000985 | +0.000982 | 0.30 % ✓
Peak ε_t | +0.000958 | +0.000984 | +0.000981 | 0.30 % ✓
Strain transfer DIC / Motor | 0.077 | 0.079 | 0.079 | 0.30 % ✓
Force relaxation in hold | 2.5 % | 2.7 % | 2.0 % | —
o Sign convention — verified — All four signals went POSITIVE together in every cycle. ✓ (Opposite of V1 — see Slide 3 for cause.)
o T6 vs T7 (post-shakedown pair): Force within 0.29 %, ε_c within 0.0 % → Best DIC repeatability on record across all campaigns.
o Shakedown completed by T5; cycles T5/T6/T7 form a stable hysteresis loop with ε_c converged near 0.00098. Earlier cycles (T0 ε_c = 0.00188, T3 ε_c = 0.00110, T4 ε_c = 0.00101) burned off the seating transient.
o Peak ε_c definition (Method a): instantaneous max |ε_c| during the displacement-hold phase. See Slide 5 for full methodology comparison.
113

## Slide 2
JÖNKÖPING UNIVERSITY
PHASE 8.6.4: VALIDATION TEST V3
Test Cases — 8.6.4 — Compression test
Results & Analysis
Same PLA specimen as V2. Five measurement cycles at −1.0 mm crosshead travel (compression), preceded by C0 conditioning. Performed AFTER Move Up/Down button fix in main.py:1990-2039; direction physically verified before acquisition.
Run | Travel | Wait before | Preload at start | Note
T0 | −1.0 mm | (after C0) | −215 N residual | Includes seating transient
T3 | −1.0 mm | ~1.5 min | Re-tared | Early shakedown
T4 | −1.0 mm | ~1.8 min | Retained | Mid shakedown
T5 | −1.0 mm | ~2.0 min | Retained | Steady-state
T6 | −1.0 mm | ~2.0 min | Retained | Final steady-state
Metric | T4 (steady) | T5 (steady) | T6 (steady) | ¦T5 vs T6¦ deviation
Peak force | −336.4 N | −333.6 N | −333.3 N | 0.07 % ✓
Peak stress | −4.205 MPa | −4.169 MPa | −4.166 MPa | 0.07 % ✓
Motor_Strain at peak | −0.10000 | −0.10000 | −0.10000 | <0.01 % ✓
ΔL_px | −7.5 px | −7.3 px | −7.2 px | 1.4 % ✓
Peak ε_c | −0.006867 | −0.006824 | −0.006750 | 1.08 % ✓
Peak ε_t | −0.006891 | −0.006847 | −0.006773 | 1.08 % ✓
Strain transfer DIC / Motor | 0.549 | 0.546 | 0.540 | 1.08 % ✓
Force relaxation in hold | 0.5 % | 0.5 % | 0.4 % | —
o Sign convention — verified — All four signals went NEGATIVE together in every cycle. ✓ (Direction verified physically post-button-fix.)
o T5 vs T6 (post-shakedown pair): Force within 0.07 %, ε_c within 1.33 % → Best Force repeatability on record.
o Hold relaxation only 0.4–0.5 % (vs tension 2.0–2.7 %) → compression engages grips firmly with no slip path. Shakedown completed by T4; earlier cycles T0/T3 (peaks −338.6 / −334.5 N) burned off seating.
o Peak ε_c definition (Method a): instantaneous max |ε_c| during the displacement-hold phase. See Slide 5 for full methodology comparison.
114

## Slide 3
JÖNKÖPING UNIVERSITY
PHASE 8.6.4: V2/V3 — SIGN CONVENTION RE-VERIFICATION
8.6.4 — Sign convention (V2/V3, corrected)
8.6.4 — Tension vs Compression per-mm
Direction | Position | Force | Motor_Strain | L_px | DIC ε_c
Tension (V2) | + | + | + | increases | +
Compression (V3) | − | − | − | decreases | −
o Change from V1: V1 reported Tension = negative / Compression = positive. V2/V3 show the opposite.
o Reason: main.py:1990-2039 contained inverted step-sign assignments for Move Up / Move Down (commented "swapped to match physical direction"). The buttons sent the motor in the opposite direction from their label.
o Consequence: V1 data tagged "Tension" was physically compression (Position −, F −); V1 data tagged "Compression" was physically tension (Position +, F +).
o Fix: main.py edited 2026-06-04 mid-session to swap the signs back. V3 was acquired with the fix in place and the direction was physically observed before testing.
Quantity per mm of travel | Tension (V2 T6/T7) | Compression (V3 T5/T6) | Ratio / Note
ΔF per mm | 757 N/mm | 333 N/mm | Tension 2.27× stiffer
ΔL_px per mm | 2.0 px/mm | 7.2 px/mm | Compression 3.6× more gauge resp.
Δε_c per mm | 0.97 × 10⁻³ | 6.74 × 10⁻³ | Compression 6.9× more strain
Strain transfer DIC/Motor | 0.078 | 0.547 | Compression 7× cleaner transfer
o Why this re-interprets V1's asymmetry
o • V1 reported '4.8× stiffer in compression'. With corrected labels, the asymmetry is the SAME magnitude (~4.6×) but in TENSION.
o • Tension grips have slack/slip — most crosshead motion is absorbed before reaching the gauge → low DIC response, high apparent force-stroke stiffness.
o • Compression presses the grips together — motion reaches the gauge directly → high DIC response, lower apparent force-stroke stiffness.
o • The direction-dependent compliance is physically meaningful and consistent across V1 (re-labelled), V2, and V3.
8.6.4 PASS ✓ Sign self-consistency (both directions) • Sign symmetry on direction reversal • Force repeatability < 5 % • DIC ε_c repeatability < 5 % • Cauchy/True agreement < 1 % • Cycle closure stable
115

## Slide 4
JÖNKÖPING UNIVERSITY
PHASE 8.6.4: VALIDATION TEST V2/V3
8.6.4 — Tension vs Compression: Cauchy vs True strain (V2/V3 representative cycles)
Test | L0 (px) | L peak (px) | ΔL_px | ε_c | ε_t | ¦ε_c − ε_t¦ | Rel. diff
V2 T0 (tension) | 1677.6 | 1680.8 | +3.2 | +0.001882 | +0.001880 | 2 × 10⁻⁶ | 0.09 %
V2 T6 (tension) | 1677.6 | 1679.3 | +1.7 | +0.000985 | +0.000984 | 1 × 10⁻⁶ | 0.05 %
V2 T7 (tension) | 1677.6 | 1679.2 | +1.6 | +0.000982 | +0.000981 | 1 × 10⁻⁶ | 0.05 %
V3 T0 (compression) | 1655.3 | 1643.9 | −11.4 | −0.006882 | −0.006906 | 24 × 10⁻⁶ | 0.35 %
V3 T5 (compression) | 1655.3 | 1644.0 | −11.3 | −0.006824 | −0.006847 | 23 × 10⁻⁶ | 0.34 %
V3 T6 (compression) | 1655.3 | 1644.1 | −11.2 | −0.006750 | −0.006773 | 23 × 10⁻⁶ | 0.34 %
Sign asymmetry is theoretically correct
o Tension (V2): L > L₀ (markers spread apart); |ε_t| < |ε_c| because ln(1+x) < x for small positive x.
o Compression (V3): L < L₀ (markers close together); |ε_t| > |ε_c| because |ln(1+x)| > |x| for small negative x.
o Physically meaningful sign direction now matches textbook: stretching ↔ positive ε; squashing ↔ negative ε.
o Cauchy and True agree to < 0.4 % across every V2/V3 cycle, exactly as the analytical small-strain expansion ε_c²/2 predicts.
o For strains < 1 % (this campaign), Cauchy and True are interchangeable. They only diverge meaningfully above ~5 % strain.
o Validates math.log(L/L0) implementation at camera_manager.py:279 across both load directions.
116

## Slide 5
JÖNKÖPING UNIVERSITY
PHASE 8.6.4: PEAK ε_c EXTRACTION METHODOLOGY
How is 'Peak ε_c' computed from ~1300 data points per cycle? Three candidates compared.
Method | Definition | Captures | Caveat
(a) Max during hold | Most extreme ¦ε_c¦ value reached in the displacement-hold window | Elastic peak + transient | Single point → noise-sensitive (~±2×10⁻⁵)
(b) Mean of first 2 s | Average ε_c over first 2 s after Position settles at peak | Pre-creep elastic state | Robust but loses 'peak' semantics
(c) Mean of full hold | Average ε_c over the entire ~20–30 s hold window | Time-averaged state | Biased by viscoelastic creep during hold
V2 Tension — peak ε_c by method
Cycle | (a) Max | (b) First 2 s | (c) Full hold
T0 | +0.001882 | +0.001854 | +0.001810
T3 | +0.001096 | +0.001060 | +0.001012
T4 | +0.001009 | +0.000982 | +0.000954
T5 | +0.000958 | +0.000929 | +0.000904
T6 | +0.000985 | +0.000955 | +0.000944
T7 | +0.000982 | +0.000931 | +0.000891
¦T6 vs T7¦ dev | 0.30 % ✓ | 2.46 % | 5.65 % ⚠
V3 Compression — peak ε_c by method
Cycle | (a) Max | (b) First 2 s | (c) Full hold
T0 | −0.006882 | −0.006844 | −0.006827
T3 | −0.006879 | −0.006831 | −0.006841
T4 | −0.006867 | −0.006841 | −0.006824
T5 | −0.006824 | −0.006778 | −0.006775
T6 | −0.006750 | −0.006715 | −0.006709
— | — | — | —
¦T5 vs T6¦ dev | 1.08 % ✓ | 0.92 % ✓ | 0.98 % ✓
o Method (a) selected. Same definition as "peak Force" — instantaneous max, not creep-averaged. V2 T6↔T7 = 0.30 %; V3 T5↔T6 = 1.08 %. Both pass the 5 % criterion.
o Why not (b)? Robust, but it averages into the creep tail. For V2 tension, (b) gives 2.46 % deviation — looks worse without changing the underlying physics.
o Why not (c)? Biased by viscoelastic creep during the 20-s hold. For V2 tension, (c) gives 5.65 % — fails the 5 % criterion through hold-window choice, not specimen behaviour.
o V3 method-insensitivity: all three methods agree within ~0.2 % because compression hold-relaxation is only 0.4–0.5 %. The conclusion is robust to method choice in compression.
117